## Slide 1: Thrombin Assay

## Slide 2: Thrombin Substrate
PEPTIDE | SEQUENCE | CHARGE
Substrate | Acetyl-N-D-D--{Nle}-TPR+/GSAGAGAG- diamino-ethyl-BFL | -1
Cleaved Fragments |  | 
N-terminal | Acetyl-N-D-D--{Nle}-TPR+-O- | -1
C-terminal | H3N+-GSAGAGAG- diamino-ethyl-BFL- | +1

## Slide 3: Thrombin Assay
Collect blood in untreated blood collection tube- Serum Tube
Combine the blood and the thrombin specific substrate
  The net charge of the thrombin substrate is -1; when thrombin cleaves produces a -2 charged fragment and a +1 fluorescently labeled fragment
Add EDTA to Stop the reaction
  EDTA chelates the Calcium in the blood, and thus Thrombin is inhibited
Separate the cleaved-labeled product from the un-cleaved substrate
  An electric field is applied to achieve this

## Slide 4
Thrombin
H3N-GSAGAGAG-NH2
BFL
Ac-NDDNleTPR-O
BFL
Ac-NDDNleTPRGSAGAGAG-NH2
BLOOD SAMPLE
SUBSTRATE

## Slide 5
BFL
Ac-NDDNleTPRGSAGAGAG-NH2
BLOOD SAMPLE
SUBSTRATE
BFL
Ac-NDDNleTPRGSAGAGAG-NH2
Thrombin
EDTA
EDTA
SUBSTRATE

## Slide 6: Thrombin Assay Diagram
Reaction Tube
} Focusing Gel
Extraction Gel
Sample Well
RBC
Thrombin
Product
Un-cleaved Substrate
EDTA
Substrate+ Blood
+ EDTA
Charge changing Substrate is added to a reaction tube
Fresh untreated blood is added to the tube
EDTA is added after some time desired
A polyacrylamide gel is loaded with the tube contents

## Slide 7: Thrombin Assay Diagram
Increase Activity
Increase Reaction Time
ANODE
CATHODE
5) Electric field is applied to the gel
6) Positive charged labeled substrate products travel to the cathode, everything else is negatively charged and travel to the anode

## Slide 8: Thrombin Assay Diagram
Increase Activity
Increase Reaction Time
} Focusing Gel
Focused Product
8) After some time the labeled positively charged products get focused in the focusing gel
9) The gel is scanned in a STORM scanner, and the image is then analyzed

## Slide 9: Sample Results

## Slide 10
Chart: Fresh Blood Thrombin Assay with EDTA
| 0.0 | 3.0 | 6.0 | 9.0 | 12.0 | 15.0 | 18.0 | 21.0 | 24.0 | 27.0 | 30.0 | 40.0 | 42.0
Blood EDTA | 2581830 | 2880851.2 | 3236558 | 4296099.4 | 5708861.6 | 7016236.6 | 8809581.2 | 9290831.6 | 9817876.6 | 10278437.2 | 10386138 | 5213103 | 8968314.5
Calibration
Blood + T2 Substrate + EDTA
Time | Sample Vol (μL) | Subst Vol (μL) | EDTA (μL)
0 | 5.8 | 0.9 | 0.9
3 | 5.8 | 0.9 | 0.9
6 | 5.8 | 0.9 | 0.9
9 | 5.8 | 0.9 | 0.9
12 | 5.8 | 0.9 | 0.9
15 | 5.8 | 0.9 | 0.9
18 | 5.8 | 0.9 | 0.9
21 | 5.8 | 0.9 | 0.9
24 | 5.8 | 0.9 | 0.9
27 | 5.8 | 0.9 | 0.9
30 | 5.8 | 0.9 | 0.9
Fresh Whole Blood - Thrombin Assay- EDTA
EDTA inhibits Thrombin action, and thus the reaction stops when it is added. As the reaction time increased, so did the thrombin activity
Reaction: Blood Samples with Substrate T2 (Thrombin); the reaction was stopped with EDTA at 3mins time intervals. Two different gels were ran in this experiment, and are represented in the graph. The total volume in the reaction was of 10 uL
Electrophoresis: 6 μL Loading in 20% Polyacrylamide, 10 min.@500V

## Slide 11
HEP Blood - Thrombin Assay- EDTA
Time | Sample Vol (μL) | Subst Vol (μL) | EDTA (μL)
0 | 5.8 | 0.9 | 0.9
3 | 5.8 | 0.9 | 0.9
6 | 5.8 | 0.9 | 0.9
9 | 5.8 | 0.9 | 0.9
12 | 5.8 | 0.9 | 0.9
15 | 5.8 | 0.9 | 0.9
18 | 5.8 | 0.9 | 0.9
21 | 5.8 | 0.9 | 0.9
24 | 5.8 | 0.9 | 0.9
27 | 5.8 | 0.9 | 0.9
30 | 5.8 | 0.9 | 0.9
No signal was detected when using HEP blood, this is as expected, since HEP inhibits thrombin activity
This can be used as a negative control, since HEP only inhibits thrombin activity, and thus since no signal is detected, then other proteases are not reacting with our thrombin substrate (T2).
Reaction: Blood Samples collected in HEP tubes reacted with Substrate T2 (Thrombin); the reaction was stopped with EDTA at 3mins time intervals. Two different gels were ran in this experiment, and are represented in the graph. The total volume in the reaction was of 10 uL
Electrophoresis: 6 μL Loading in 20% Polyacrylamide, 10 min.@500V
Chart: HEP Blood Thrombin Assay with EDTA
| 0.0 | 3.0 | 6.0 | 9.0 | 12.0 | 15.0 | 18.0 | 21.0 | 24.0 | 27.0 | 30.0 | 40.0 | 42.0
HEP-Blood EDTA | 2509039.6 | 2614092 | 2610324 | 2710667.2 | 2796876 | 2810164 | 2755796.4 | 2748520.6 | 2719787.4 | 2800093.6 | 2984241.6 | 5476857.2 | 9249753
Calibration

## Slide 12
Aspirin Fresh Whole Blood- Thrombin Assay- EDTA
Time | Sample Vol (μL) | Subst Vol (μL) | EDTA (μL)
0 | 5.8 | 0.9 | 0.9
3 | 5.8 | 0.9 | 0.9
6 | 5.8 | 0.9 | 0.9
9 | 5.8 | 0.9 | 0.9
12 | 5.8 | 0.9 | 0.9
15 | 5.8 | 0.9 | 0.9
18 | 5.8 | 0.9 | 0.9
21 | 5.8 | 0.9 | 0.9
24 | 5.8 | 0.9 | 0.9
27 | 5.8 | 0.9 | 0.9
30 | 5.8 | 0.9 | 0.9
NA | Negative Control | Negative Control | Negative Control
NA | Control | Control | Control
Fresh blood was collected from subject, who took 2 aspirin pills before the blood draw (one at night, and one in the morning). The detection of signal of thrombin activity was delayed by approx. 5 min.
Reaction: Blood Samples ( containing traces of Aspirin) reacted with Substrate T2 (Thrombin); the reaction was stopped with EDTA at 3mins time intervals. Two different gels were ran in this experiment, and are represented in the graph. The total volume in the reaction was of 10 uL
Electrophoresis: 6 μL Loading in 20% Polyacrylamide, 10 min.@500V
Chart: Aspirin Fresh Blood Thrombin Assay with EDTA
| 0.0 | 3.0 | 6.0 | 9.0 | 12.0 | 15.0 | 18.0 | 21.0 | 24.0 | 29.0 | 34.0 | 40.0 | 42.0
Blood-Aspirin EDTA | 2367510 | 2440500 | 2338248.8 | 2306750 | 2205020.8 | 2582526 | 4483623.2 | 4384272 | 9236159.2 | 10849174 | 11406892.8 | 3242669.6 | 9485075

## Slide 13
Fresh Whole Blood- Thrombin Assay- EDTA (Aspirin and normal samples)
Normal Blood Sample:
Aspirin Blood Sample:
0 Reaction Time (min) 35
0 Reaction Time (min) 37
Buffer
Buffer

## Slide 14
Lane | [Thrombin] nM
1 | 600
3 | 500
5 | 400
7 | 300
9 | 200
11 | 100
13 | 30
14 | Neg Control
15 | Neg Control
Calibration Curves: Citrate Blood +Thrombin
Citrated blood was added with Thrombin substrate (T2), and thrombin enzyme
As the thrombin enzyme increased so did the fluorescent signal
Reaction: Blood Samples with Substrate T2 (Thrombin) and thrombin enzyme ; the reactants were added into a reaction tube, and the reaction was allowed for 30 mins. The total volume in the reaction was 15uL.
Electrophoresis: 6 μL Loading in 20% Polyacrylamide, 10 min.@500V
Chart: Thrombin Calibration Curve in Citrate Blood
| 600.0 | 500.0 | 400.0 | 300.0 | 200.0 | 100.0 | 30.0
 | 13179664 | 10066469 | 8979616.8 | 8806589.2 | 7021682 | 6064664.8 | 5380849

## Slide 15: Bacterial Proteases

## Slide 16: OmpT
Outer membrane Bacterial Protease expressed by Escherichia coli
Structure:
  10 stranded antiparallel β- Barrels , which protrudes out of the lipid bilayer
Active Site:
  Proteolytic activity in the extracellular part of the membrane
  This extracellular part of OMPT contains a large negatively charge groove: thus it has preference for Positively charged Residues

## Slide 17: OmpT Structure
Electrostatic Surface Potential
Overall Structure of OmpT
Negative Charged Residues
Catalytic Residues
EMBO J. 2001 September 17; 20(18): 5033–5039.

## Slide 18: OmpT Structure
EMBO J. 2001 September 17; 20(18): 5033–5039.

## Slide 19: OmpT Substrate Design
The substrate needs to meet the following requirements:
  Specificity
  Sensitivity
  The whole substrate needs to have a net NEGATIVE CHARGE, the labeled fragment caused by the enzymatic cleavage of the substrate needs to have POSITIVE CHARGE
OmpT-Substrate Binding:
  Amino acids in the substrate within six residue window contribute to the binding of it to OmpT:
  P1: Arg requirement
  P1’: Lys, Gly, Val , Arg
  P2: Val or Al
  P3 and P4: Trp or Arg

## Slide 20: OmpT Substrate Design
Strategy
  Identify OmpT specific peptide that is most negatively charged or close to neutral
  Modify the ends of the peptide with charged amino acid residues
  Determine the net charge of the entire substrate: must be negative. And also determine charge of the cleaved fractions: one of them must be positive (c-end)
  Bind fluorophore to the positively charged fraction terminal

## Slide 21: OmpT Substrate Design
Substrate Sequence of Interest
  Determined by the following techniques
    In situ cleavage of phage that display protease-susceptible peptides by E coli expressing OmpT
    In vitro cleavage of phage-displayed peptides using purified enzyme
  McCarter’s group narrowed it down to the following after testing all combinations in their library:

## Slide 22: OmpT Substrate Design
  Best sequence for charge changing substrate:
  Ac-WGGK(+)YR(+) / R(+)AWGTI-NH2
    The peptide shown above has a net charge of +3. At least 4 negatively charged residues need to be added to the sequence, making the net charge of the substrate -1
    - Must keep the peptide as short as possible to avoid complications in structure
    Notice the synthetic substrate sequence is 106 fold higher than peptide that corresponds to cleavage site of human plasminogen
    Suggested Sequence:
    Use Asp(D) as the negatively charged residue: less bulky than Glu(E)
    Add most of the negative residues to the N-terminus, since this will be the unlabeled end.
      This will make the label end of the substrate more positive, and the un-label more negative  Better separation

## Slide 23: OmpT Substrate Design
    Suggested Sequence:
  Acetyl-W-GD(-)D(-)GD(-) K(+) YR(+) / R(+)AWGD(-)TI(diamino-ethyl BFL)-Ac
PEPTIDE | SEQUENCE | CHARGE
Substrate | Acetyl-W-GD(-)D(-)GD(-) K(+) YR(+) / R(+)AWGD(-)TI(diamino-ethyl BFL)-Ac | -1
Cleaved Fragments |  | 
N-terminal | Acetyl-W-GD(-)D(-)GD(-) K(+) YR(+) -O- | -2
C-terminal | H3N+-R(+)AWGD(-)TI(diamino-ethyl BFL)-Ac | +1

## Slide 24: OmpT Assay
Similar to the thrombin assay, with the exception of a different substrate that targets OmpT.
Sample with OmpT enzyme
OmpT Specific substrate
Cleaved substrate

## Slide 25: Pla
Yersinia pestis surface protease
An agent of plague and has been recognized as one of the most devastating, epidemic-causing bacteria experienced by mankind
9.5-kb plasmid pPCP expresses plague plasminogen activator (Pla)
Responsible for fibrinolytic and coagulase activities

## Slide 26: Pla Substrate Design
Design strategy is very similar to the OmpT substrate
DABCYL-Arg(+) / Arg(+)-Ile-Asn-Arg(+)-Glu(-)(EDANS)-NH2
DABCYL is quencher molecule, EDANS is fluorophore
+2 charge peptide sequence. Thus, we have to add -3 charge

## Slide 27: Point of Care (POC) Device Hand Held Device

## Slide 28: POC Device Sketch
Buffer Well: 1) 10- 20 μL
Electrode
Sample Well: 2) 5- 10 μL
and Lyophilized Substrate
Electrode
Channel with Filter Paper
Negatively Charged Beads
Support:
Glass or Plastic
Top View:

## Slide 29: POC Device Sketch
Electrode
Electrode
Buffer Well
Sample Well
Side View:
Wire Connection
Wire Connection
Lyophilized Reagents:
-Charge Changing Substrate, Cofactors…
Cover: glass or plastic: transparent

## Slide 30: POC Device Sketch
`
Buffer is added to the buffer gel:
    Wetting the filter paper  electrolyte connection between buffer and the sample wells
2) Sample is added in the sample well:
    Solubilizing the lyophilized reagents, and providing a connection between the sample, and buffer wells
3) The reagents, and the sample are allowed to react for 10-15 min
4) A DC electric field is applied
    Sample chamber (+) charge; Buffer chamber (-) charge
5) The positive charge fluorescently labeled product is collected in the beads
`

## Slide 31: POC Device Sketch
`
`
`
Batteries
  6) After the product is collected in the beads, imaging can be done by adding an excitation source (laser) under the device, then a filter on top, and the image is detected by a CCD camera on top.
Detector: CCD Camera
Filter
Excitation Source: Laser

## Slide 32: CATHEPSIN-S

## Slide 33: CATHEPSIN-S
Lysosomal cysteine protease
  Involve in multiple diseases: cancer, diabetes, cardiovascular diseases…
and…

## Slide 34: CATHEPSIN-S
GRWHTVG//LRWE-Lys(Dnp)-DArg- NH2
From the sequence above, single amino-acid were replaced, the result is shown:
The polypeptide seemed to be more sensitive to changes in P1’, P2,P1, and P3

## Slide 35: CATHEPSIN-S
Broad substrate specificity  Specific substrate is not known, but from the sequence shown before, the substrate seems to prefer the followings:
Side Chain |  | P-2 | P-1 | P-1’ | P-3’
Hydrophobic | Aromatic Aliphatic Methionine Proline | +(only Phe) +(not Ile) + - | - + - - | - + - - | + + + Not tested
Neutral | Amide Small | - - | + + | - - | 0 -
Basic |  | - - | + + | 0 - | - -
Acidic |  | - | - | - | -

## Slide 36: CATHEPSIN-S

## Slide 37: CATHEPSIN-S
Comparing Substrates
Ac-N-GD- PVG // LTA GAGK(BFL)-NH2 : Substrate for MMP2-9
Ac-GRWH-PMG // LPW ELys(Dnp)-DArg- NH2 : Substrate with higher specificity for CatS when compared to CatL, and CatB
The substrates are similar in the recognition sequence, so there might be non-specific cleavage of the substrate by MMP2/9 instead of Cathepsin S  might be better to sacrify the specificity in comparison to CatB and CatL than that due to MMP2/9 (elevated in the blood-from the diabetes studies)
Might be better to go with: Ac-GRWH-TVG // LR WE-Lys(Dnp)-DArg- NH2

## Slide 38: CATHEPSIN-S
Substrate 1:
Ac-GR(+)WH-TV G // L R(+) WE-dR(+)-K(+)(BFL) NH2